## Trigger Pull Data Version 2
Date | Action | Revolver | [6 empty] | Pistol 1 | Pistol 2 | Rifle 1 | Rifle 2 | Shotgun 1 | Shotgun 2 | Firearm | Overall Stats
|  | SA | DA - C1 | DA - C2 | DA - C3 | DA - C4 | DA - C5 | DA - C6 | SA | SA | SA | SA | SA | SA
[16 empty] | Means | Stdevs | Relative sd
Day 1 |  | 10.23 | 14.25 | 14.5 | 14.5 | 14.25 | 14.5 | 15 | 9.948 | 5.75 | 5.187 | 6.75 | 3.855 | 5.67 |  | lbf | lbf
|  | 11.08 | 15.25 | 15.5 | 16 | 15.75 | 15.5 | 16 | 10.4 | 6 | 4.848 | 6.75 | 4.252 | 5.5 | Revolver | 10.819 | 0.80589 | 0.074488
|  | 10.13 | 15.25 | 15.25 | 15.25 | 15 | 15.5 | 15.75 | 7.69 | 6.75 | 4.38 | 7.083 | 3.86 | 5.75 | C1 | 15.175 | 0.4091 | 0.026959
|  | 10.16 | 15.5 | 15 | 15 | 15 | 14.75 | 15.25 | 7.02 | 5.75 | 4.97 | 7.25 | 4.04 | 5.67 | C2 | 14.725 | 0.43221 | 0.029352
|  | 10.88 | 15.5 | 14.5 | 15 | 14.5 | 15 | 15.25 | 8.21 | 6.5 | 3.91 | 7.25 | 4.87 | 5.92 | C3 | 14.85 | 0.50277 | 0.033857
Day 2 |  | 10.63 | 15.5 | 14.25 | 14.75 | 15 | 14.75 | 15.75 | 7.87 | 6.5 | 5.21 | 7.25 | 4.78 | 5.58 | C4 | 14.775 | 0.46323 | 0.031352
|  | 10.49 | 14.75 | 14.5 | 14.25 | 14.5 | 14.75 | 15 | 7.82 | 6.75 | 5.23 | 7.25 | 3.77 | 5.33 | C5 | 14.85 | 0.41164 | 0.02772
|  | 10.85 | 15.25 | 15 | 14.75 | 14.5 | 14.5 | 15.5 | 6.99 | 6.25 | 4.38 | 7.083 | 4.24 | 5.83 | C6 | 15.325 | 0.4091 | 0.026695
|  | 12.91 | 15.5 | 14.25 | 14.5 | 15 | 15 | 15 | 9.04 | 6.5 | 5.93 | 6.913 | 4.73 | 5.58
|  | 10.83 | 15 | 14.5 | 14.5 | 14.25 | 14.25 | 14.75 | 8.78 | 6.75 | 4.73 | 7.083 | 4.45 | 5.42 | Pistol 1 | 8.3768 | 1.1545 | 0.13782
[15 empty] | Pistol 2 | 6.35 | 0.39441 | 0.062111
[15 empty] | Rifle 1 | 4.8775 | 0.56771 | 0.11639
[15 empty] | Rifle 2 | 7.0662 | 0.19984 | 0.028281
[15 empty] | Shotgun 1 | 4.2847 | 0.40912 | 0.095483
[15 empty] | Shotgun 2 | 5.625 | 0.1818 | 0.032319
We use the largest std. dev. for each firearm/action type in the process study.
Enter these values on the Measurement Uncertainty Estimation Forms at cell C8.
Total Number of Observations for all firearms:
130
Number of Repetitions of the Same Quantity:
10
Data and calculations adapted from results by:
Unified Metropolitan Forensic Crime Laboratory
Example Spreadsheet for Calculating Uncertainty for Trigger Pull Measurements
July 12, 2019
The analysis begins below on Sheet 1 with simulated data involving measurements taken under the scenario described in Clause 4.5 of the text. This includes 10 measurements each for 2 rifles, 2 shotguns, 2 pistols, and one revolver operating in single action. The scenario also includes 10 measurements for each cylinder position on a 6-shot revolver operating in double action. All measurements in this example are assumed to be taken by a single expert. Each measurement shown below is assumed to be an average of three successive trigger pulls. Altogether 130 measurements are shown below, representing 390 trigger pulls. If more than one expert is used in the study, the workload of 130 measurements can be shared. The calculations here yield a mean value for the trigger pull of each firearm and a Type A standard uncertainty (standard deviation). The largest standard deviation is recorded for each firearm type. The standard deviations are then carried onto subsequent spreadsheets and combined with other sources of uncertainty.

## Levene's Modified Test
| Revolver |  | Pistol 1 | Rifle 1 | Shotgun 1
| SA | DA - C3 | SA | SA | SA
| 10.23 | 14.5 | 9.948 | 5.187 | 3.855
| 11.08 | 16 | 10.4 | 4.848 | 4.252
| 10.13 | 15.25 | 7.69 | 4.38 | 3.86
| 10.16 | 15 | 7.02 | 4.97 | 4.04
| 10.88 | 15 | 8.21 | 3.91 | 4.87
| 10.63 | 14.75 | 7.87 | 5.21 | 4.78
| 10.49 | 14.25 | 7.82 | 5.23 | 3.77
| 10.85 | 14.75 | 6.99 | 4.38 | 4.24
| 12.91 | 14.5 | 9.04 | 5.93 | 4.73
| 10.83 | 14.5 | 8.78 | 4.73 | 4.45
Average | 10.819 | 14.85 | 8.3768 | 4.8775 | 4.2847
Median | 10.73 | 14.75 | 8.04 | 4.909 | 4.246
| Revolver |  | Pistol 1 | Rifle 1 | Shotgun 1
| SA | DA - C3 | SA | SA | SA
Absolute value of the difference between each average value and the median | 0.5 | 0.25 | 1.908 | 0.278 | 0.391
| 0.35 | 1.25 | 2.36 | 0.061 | 0.006
| 0.6 | 0.5 | 0.35 | 0.529 | 0.386
| 0.57 | 0.25 | 1.02 | 0.061 | 0.206
| 0.15 | 0.25 | 0.17 | 0.999 | 0.624
| 0.1 | 0 | 0.17 | 0.301 | 0.534
| 0.24 | 0.5 | 0.22 | 0.321 | 0.476
| 0.12 | 0 | 1.05 | 0.529 | 0.006
| 2.18 | 0.25 | 1 | 1.021 | 0.484
| 0.1 | 0.25 | 0.74 | 0.179 | 0.204
Anova: Single Factor
SUMMARY
Groups | Count | Sum | Average | Variance
Column 1 | 10 | 4.91 | 0.491 | 0.39039
Column 2 | 10 | 3.5 | 0.35 | 0.12778
Column 3 | 10 | 8.988 | 0.8988 | 0.56121
Column 4 | 10 | 4.279 | 0.4279 | 0.11995
Column 5 | 10 | 3.317 | 0.3317 | 0.046791
ANOVA
Source of Variation | SS | df | MS | F | P-value | F crit
Between Groups | 2.1515 | 4 | 0.53786 | 2.1582 | 0.089106 | 2.0742
Within Groups | 11.215 | 45 | 0.24922
Total | 13.3665 | 49
F is larger than F crit at the 10 % level.
Therefore, the standard deviations of the trigger pull measurements
for different firearm types are considered to be significantly different
using Levene's Modified Test.
In the Data Analysis Toolpack add-in, select the one-way ANOVA. For the input range, select all cells containing the absolute value of the difference between each average value and the median. Grouped by column. Select 0.1 alpha.

## Revolver in SA
Measurement Uncertainty Estimation Form
Measurement: |  |  | Trigger Pull Force of a Revolver in Single Action
Range of measurement values: |  |  | 10.13 lbsf - 12.91 lbsf
Procedure name and revision: |  |  | Example Uncertainty Budget
Estimation prepared by: |  |  | Task Group on Uncertainty of Measurement |  |  |  |  | Date Prepared: |  | 2019-07-04
Attention! See explanatory notes on row 7 and column A.
Line Item | Uncertainty Component | Value | Units | Type | Statistical Distribution | Divisor | Degrees Freedom (n-1) | Standard Uncertainty (lbsf) | Component Contribution %
1 | Measurement Process Reproducibility (from page 1) | 0.806 | lbsf | A | normal | 1 | 19 | 0.806 | 99.3991
2 | Calibration Certificate for Force Gauge - ± 1 % | 0.108 | lbsf | B | rectangular | 1.73 | ∞ | 0.062428 | 0.5963
3 | Thermal Dependence of Force Gauge. Assume ± 5 °F uncertainty in temperature. | 0.0095 | lbsf | B | rectangular | 1.73 | ∞ | 0.0054913 | 0.0046139
4 | Other uncertainty sources, etc.
[8 empty] | Sum | 100
| Combined Standard Unc | uc | [5 empty] | 0.80843
| Expanded Unc | U (k=2) | [5 empty] | 1.6169
| Expanded Unc | U (k=3) | [5 empty] | 2.4253
| Reported Uncertainty: | 1.6169 |  | k=2
| Reported Uncertainty: | 2.4253 |  | k=3
NOTE: Regardless of the number of digits that are showing in a cell, Excel carries the maximum number of significant figures in the background and will use the entire number for further calculations
Note on A8: The first component of uncertainty is an estimate of the standard deviation of the measurements by the observers in the study. It is copied from cell T20 shown on sheet one. -Chris M
Note on A9: The force gauge should have an uncertainty obtainable from the vendor or from a calibration lab. For simplicity, we assume here an uncertainty of ± 1 % maximum with a rectangular distribution. With a measured average value of 10.819 lbsf shown on sheet 1, the estimated maximum uncertainty is 0.108 lbsf. This component of uncertainty may also be quoted by the vendor as a fixed amount (example - 0.1 lbsf). -Chris M
Note on A10: Force measuring instruments should have some temperature dependence, particularly if they have any spring components. The last worksheet contains a rough estimate of possible error in a spring constant, yielding ± 0.088 % (outer limit of error) for a temperature uncertainty of ± 5 °F. With a measured value of 10.819 lbsf, the maximum uncertainty is ± 0.0095 lbsf. -Chris M
Note on C7: Column C shows the estimated uncertainty values from each source. This is either derived from data showing variation of the source component or is derived from our understanding and modelling of the possible bias or variation. -LJFarrell
Note on C17: Limit to at most two significant figures unless there is a documented rationale for reporting additional significant figures. The laboratory procedure for estimating measurement uncertainty must include the process for rounding the expanded uncertainty. -LJFarrell
Note on D7: Column D: The source components in this template are all converted to lbsf. If one source uncertainty component has different units from the others, such as newtons or %, the values need to be converted into equivalent values in lbsf, and the converted values need to be recorded in column C for combination with the other uncertainty components. -LJFarrell
Note on E7: Column E: If the uncertainty value in column C is derived by statistical methods yielding, say, a standard deviation, then the uncertainty is Type A. If the value is derived by other means such as from manufacturer specifications or from a calculation of possible bias, then the uncertainty is Type B. -LJFarrell
Note on F7: Columns F and G: If Column E is Type A, then the distribution of expected values is assumed to be "normal," also called Gaussian. If Column E is Type B, other distributions can be assumed. If a full range of possible values is estimated in column C , then a rectangular distribution of possible values is usually assumed. If so, then the standard uncertainty is equal to that full range divided by 2√3. If only one side of the range of possible values is shown in Column C, as in cells C9 and C10 here, then the divisor for a rectangular distribution is √3. -LJFarrell
Note on G7: Columns F and G: If Column E is Type A, then the distribution of expected values is assumed to be "normal," also called Gaussian. If Column E is Type B, other distributions can be assumed. If a full range of possible values is estimated in column C , then a rectangular distribution of possible values is usually assumed. If so, then the standard uncertainty is equal to that full range divided by 2√3. If only one side of the range of possible values is shown in Column C, as in cells C9 and C10 here, then the divisor for a rectangular distribution is √3. -LJFarrell
Note on H7: Column H: The more data points you have, the more confident you can be that a statistical calculation of 95 % confidence intervals is accurate. The degrees of freedom may be used to estimate confidence intervals. For type B components calculated by non-statistical methods, people usually calculate a conservative worst case estimate of error and assume infinite d.o.f. -Chris M
Note on I7: Column I = Column C / Column G. This operation converts the stated uncertainty values in Column C into standard uncertainties. Now we are in a position to combine apples with apples. The quadratic sum is the combined standard uncertainty, and it appears in cell I14 here. -LJFarrell
Note on J7: Column J: See the formulas in the cells in this column. The largest component contributes the most to the resulting uncertainty. -Chris M

## Revolver in DA
Measurement Uncertainty Estimation Form
Measurement: |  |  | Trigger Pull Force of a Revolver in DA
Range of measurement values: |  |  | 14.25 lbsf - 16.00 lbsf
Procedure name and revision: |  |  | Example Uncertainty Budget
Estimation prepared by: |  |  | Task Group on Uncertainty of Measurement |  |  |  |  | Date Prepared: |  | 2019-07-04
Attention! See explanatory notes on row 7 and column A.
Line Item | Uncertainty Component | Value | Units | Type | Statistical Distribution | Divisor | Degrees Freedom (n-1) | Standard Uncertainty (lbsf) | Component Contribution %
1 | Measurement Process Reproducibility (from page 1) | 0.503 | lbsf | A | normal | 1 | 19 | 0.503 | 97.1306
2 | Calibration Certificate for Force Gauge - ± 1 % | 0.149 | lbsf | B | rectangular | 1.73 | ∞ | 0.086127 | 2.8477
3 | Thermal Dependence of Force Gauge. Assume ± 5°F uncertainty in temperature. | 0.013 | lbsf | B | rectangular | 1.73 | ∞ | 0.0075145 | 0.021678
4 | Other uncertainty sources, etc.
[8 empty] | Sum | 100
| Combined Standard Unc | uc | [5 empty] | 0.51038
| Expanded Unc | U (k=2) | [5 empty] | 1.0208
| Expanded Unc | U (k=3) | [5 empty] | 1.5311
| Reported Uncertainty: | 1.0208 |  | k=2
| Reported Uncertainty: | 1.5311 |  | k=3
NOTE: Regardless of the number of digits that are showing in a cell, Excel carries the maximum number of significant figures in the background and will use the entire number for further calculations
Note on A8: The first component of uncertainty is an estimate of the standard deviation of the measurements by the observers in the study. It is copied from cell T23 shown on sheet one. -Chris M
Note on A9: The force gauge should have an uncertainty obtainable from the vendor or from a calibration lab. For simplicity, we assume here a maximum uncertainty of ± 1 % with a rectangular distribution. This component of uncertainty may also be quoted by the vendor as a fixed amount (example - 0.1 lbsf). With a measured average value of 14.850 lbsf shown on page 1, the estimated maximum uncertainty is 0.149 lbsf. -Chris M
Note on A10: It seems that force measuring instruments should have some temperature dependence, particularly if they have any spring components. The last worksheet contains a rough estimate of possible error in a spring constant, yielding ± 0.088 % (outer limit of error) for a temperature uncertainty of ± 5 °F. With a measured value of 14.850 lbsf, the maximum uncertainty is ± 0.0130 lbsf.last -Chris M
Note on A11: other sources, etc. -Chris M
Note on B7: Column B shows the different sources of uncertainty that are considered significant. -Chris M
Note on C7: Column C shows the estimated uncertainty values from each source. This is either derived from data showing variation of the source component or is derived from our understanding and modelling of the possible bias or variation. -LJFarrell
Note on C17: Limit to at most two significant figures unless there is a documented rationale for reporting additional significant figures. The laboratory procedure for estimating measurement uncertainty must include the process for rounding the expanded uncertainty. -LJFarrell
Note on D7: Column D: The source components in this template are all converted to lbsf. If one source uncertainty component has different units from the others, such as newtons or %, the values need to be converted into equivalent values in lbsf, and the converted values need to be recorded in column C for combination with the other uncertainty components. -LJFarrell
Note on E7: Column E: If the uncertainty value in column C is derived by statistical methods yielding, say, a standard deviation, then the uncertainty is Type A. If the value is derived by other means such as from manufacturer specifications or from a calculation of possible bias, then the uncertainty is Type B. -LJFarrell
Note on F7: Columns F and G: If Column E is Type A, then the distribution of expected values is assumed to be "normal," also called Gaussian. If Column E is Type B, other distributions can be assumed. If a full range of possible values is estimated in column C , then a rectangular distribution of possible values is usually assumed. If so, then the standard uncertainty is equal to that full range divided by 2√3. If only one side of the range of possible values is shown in Column C, as in cells C9 and C10 here, then the divisor for a rectangular distribution is √3. -LJFarrell
Note on G7: Columns F and G: If Column E is Type A, then the distribution of expected values is assumed to be "normal," also called Gaussian. If Column E is Type B, other distributions can be assumed. If a full range of possible values is estimated in column C , then a rectangular distribution of possible values is usually assumed. If so, then the standard uncertainty is equal to that full range divided by 2√3. If only one side of the range of possible values is shown in Column C, as in cells C9 and C10 here, then the divisor for a rectangular distribution is √3. -LJFarrell
Note on H7: Column H: The more data points you have, the more confident you can be that a statistical calculation of 95 % confidence intervals is accurate. The degrees of freedom may be used to estimate confidence intervals. For type B components calculated by non-statistical methods, people usually calculate a conservative worst case estimate of error and assume infinite d.o.f. -Chris M
Note on I7: Column I = Column C / Column G. This operation converts the stated uncertainty values in Column C into standard uncertainties. Now we are in a position to combine apples with apples. The quadratic sum is the combined standard uncertainty, and it appears in cell I14 here. -LJFarrell
Note on J7: Column J: See the formulas in the cells in this column. The largest component contributes the most to the resulting uncertainty. -Chris M

## Pistol
Measurement Uncertainty Estimation Form
Measurement: |  |  | Trigger Pull Force of Pistols
Range of measurement values: |  |  | 5.75 lbsf - 10.40 lbsf
Procedure name and revision: |  |  | Example Uncertainty Budget
Estimation prepared by: |  |  | Task Group on Uncertainty of Measurement |  |  |  |  | Date Prepared: |  | 2019-07-04
Attention! See explanatory notes on row 7 and column A.
Line Item | Uncertainty Component | Value | Units | Type | Statistical Distribution | Divisor | Degrees Freedom (n-1) | Relative Standard Uncertainty (%) | Component Contribution %
1 | Measurement Process Reproducibility (from page 1) | 1.154 | lbsf | A | normal | 1 | 19 | 1.154 | 99.8219
2 | Calibration Certificate for Force Gauge - ± 1 % | 0.084 | lbsf | B | rectangular | 1.73 | ∞ | 0.048555 | 0.17672
3 | Thermal Dependence of Force Gauge. Assume ± 5 °F uncertainty in temperature. | 0.0073 | lbsf | B | rectangular | 1.73 | ∞ | 0.0042197 | 0.0013347
4 | Other uncertainty sources, etc.
[8 empty] | Sum | 100
| Combined Standard Unc | uc | [5 empty] | 1.155
| Expanded Unc | U (k=2) | [5 empty] | 2.3101
| Expanded Unc | U (k=3) | [5 empty] | 3.4651
| Reported Uncertainty: | 2.3101 |  | k=2
| Reported Uncertainty: | 3.4651 |  | k=3
NOTE: Regardless of the number of digits that are showing in a cell, Excel carries the maximum number of significant figures in the background and will use the entire number for further calculations.
Note on A8: The first component of uncertainty is an estimate of the standard deviation of the measurements by the observers in the study. It is copied from cell T28 shown on sheet one. -Chris M
Note on A9: The force gauge should have an uncertainty obtainable from the vendor or from a calibration lab. For simplicity, we assume here a maximum uncertainty of ± 1 % with a rectangular distribution. This component of uncertainty may also be quoted by the vendor as a fixed amount (example - 0.1 lbsf). With a measured average value of 8.377 lbsf shown on page 1, the estimated maximum uncertainty is 0.084 lbsf. -Chris M
Note on A10: It seems that force measuring instruments should have some temperature dependence, particularly if they have any spring components. The last worksheet contains a rough estimate of possible error in a spring constant, yielding ± 0.088 % (outer limit of error) for a temperature uncertainty of ± 5 °F. With a measured value of 8.377 lbsf, the maximum uncertainty is ± 0.0073 lbsf. -Chris M
Note on A11: other sources, etc. -Chris M
Note on B7: Column B shows the different sources of uncertainty that are considered significant. -Chris M
Note on C7: Column C shows the estimated uncertainty values from each source. This is either derived from data showing variation of the source component or is derived from our understanding and modelling of the possible bias or variation. -LJFarrell
Note on C17: Limit to at most two significant figures unless there is a documented rationale for reporting additional significant figures. The laboratory procedure for estimating measurement uncertainty must include the process for rounding the expanded uncertainty. -LJFarrell
Note on D7: Column D: The source components in this template are all converted to lbsf. If one source uncertainty component has different units from the others, such as newtons or %, the values need to be converted into equivalent values in lbsf, and the converted values need to be recorded in column C for combination with the other uncertainty components. -LJFarrell
Note on E7: Column E: If the uncertainty value in column C is derived by statistical methods yielding, say, a standard deviation, then the uncertainty is Type A. If the value is derived by other means such as from manufacturer specifications or from a calculation of possible bias, then the uncertainty is Type B. -LJFarrell
Note on F7: Columns F and G: If Column E is Type A, then the distribution of expected values is assumed to be "normal," also called Gaussian. If Column E is Type B, other distributions can be assumed. If a full range of possible values is estimated in column C , then a rectangular distribution of possible values is usually assumed. If so, then the standard uncertainty is equal to that full range divided by 2√3. If only one side of the range of possible values is shown in Column C, as in calls C9 and C10, then the divisor for a rectangular distribution is √3. -LJFarrell
Note on G7: Columns F and G: If Column E is Type A, then the distribution of expected values is assumed to be "normal," also called Gaussian. If Column E is Type B, other distributions can be assumed. If a full range of possible values is estimated in column C , then a rectangular distribution of possible values is usually assumed. If so, then the standard uncertainty is equal to that full range divided by 2√3. If only one side of the range of possible values is shown in Column C, as in calls C9 and C10 here, then the divisor for a rectangular distribution is √3. -LJFarrell
Note on H7: Column H: The more data points you have, the more confident you can be that a statistical calculation of 95 % confidence intervals is accurate. The degrees of freedom may be used to estimate confidence intervals. For type B components calculated by non-statistical methods, people usually calculate a conservative worst case estimate of error and assume infinite d.o.f. -Chris M
Note on I7: Column I = Column C / Column G. This procedure converts the stated uncertainty values in Column C into standard uncertainties. Now we are in a position to combine apples with apples. The quadratic sum is the combined standard uncertainty, and it appears in cell I14 here. -LJFarrell
Note on J7: Column J: See the formulas in the cells in this column. The largest component contributes the most to the resulting uncertainty. -Chris M

## Rifle
Measurement Uncertainty Estimation Form
Measurement: |  |  | Trigger Pull Force of Rifle 1
Range of measurement values: |  |  | 3.91 lbsf - 7.25 lbsf
Procedure name and revision: |  |  | Example Uncertainty Budget
Estimation prepared by: |  |  | Task Group on Uncertainty of Measurement |  |  |  |  | Date Prepared: |  | 2019-07-04
Attention! See explanatory notes on row 7 and column A.
Line Item | Uncertainty Component | Value | Units | Type | Statistical Distribution | Divisor | Degrees Freedom (n-1) | Relative Standard Uncertainty (%) | Component Contribution %
1 | Measurement Process Reproducibility (from page 1) | 0.568 | lbsf | A | normal | 1 | 19 | 0.568 | 99.7501
2 | Calibration Certificate for Force Gauge - ± 1 % | 0.049 | lbsf | B | rectangular | 1.73 | ∞ | 0.028324 | 0.24804
3 | Thermal Dependence of Force Gauge. Assume ± 5°F uncertainty in temperature. | 0.0043 | lbsf | B | rectangular | 1.73 | ∞ | 0.0024855 | 0.0019101
4 | Other uncertainty sources, etc.
[8 empty] | Sum | 100
| Combined Standard Unc | uc | [5 empty] | 0.56871
| Expanded Unc | U (k=2) | [5 empty] | 1.1374
| Expanded Unc | U (k=3) | [5 empty] | 1.7061
| Reported Uncertainty: | 1.1374 |  | k=2
| Reported Uncertainty: | 1.7061 |  | k=3
NOTE: Regardless of the number of digits that are showing in a cell, Excel carries the maximum number of significant figures in the background and will use the entire number for further calculations
Note on A8: The first component of uncertainty is an estimate of the standard deviation of the measurements by the observers in the study. It is copied from cell T32 shown on sheet one. -Chris M
Note on A9: The force gauge should have an uncertainty obtainable from the vendor or from a calibration lab. For simplicity, we assume here a maximum uncertainty of ± 1 % with a rectangular distribution. This component of uncertainty may also be quoted by the vendor as a fixed amount (example - 0.1 lbsf). With a measured average value of 4.877 lbsf shown on page 1, the estimated maximum uncertainty is 0.049 lbsf. -Chris M
Note on A10: It seems that force measuring instruments should have some temperature dependence, particularly if they have any spring components. The last worksheet contains a rough estimate of possible error in a spring constant, yielding ± 0.088 % (outer limit of error) for a temperature uncertainty of ± 5 °F. With a measured value of 4.877 lbsf, the maximum uncertainty is ± 0.0043 lbsf. -Chris M
Note on A11: other sources, etc. -Chris M
Note on B7: Column B shows the different sources of uncertainty that are considered significant. -Chris M
Note on C7: Column C shows the estimated uncertainty values from each source. This is either derived from data showing variation of the source component or is derived from our understanding and modelling of the possible bias or variation. -LJFarrell
Note on C17: Limit to at most two significant figures unless there is a documented rationale for reporting additional significant figures. The laboratory procedure for estimating measurement uncertainty must include the process for rounding the expanded uncertainty. -LJFarrell
Note on D7: Column D: The source components in this template are all converted to lbsf. If one source uncertainty component has different units from the others, such as newtons or %, the values need to be converted into equivalent values in lbsf, and the converted values need to be recorded in column C for combination with the other uncertainty components. -LJFarrell
Note on E7: Column E: If the uncertainty value in column C is derived by statistical methods yielding, say, a standard deviation, then the uncertainty is Type A. If the value is derived by other means such as from manufacturer specifications or from a calculation of possible bias, then the uncertainty is Type B. -LJFarrell
Note on F7: Columns F and G: If Column E is Type A, then the distribution of expected values is assumed to be "normal," also called Gaussian. If Column E is Type B, other distributions can be assumed. If a full range of possible values is estimated in column C , then a rectangular distribution of possible values is usually assumed. If so, then the standard uncertainty is equal to that full range divided by 2√3. If only one side of the range of possible values is shown in Column C, as in cells C9 and C10 here, then the divisor for a rectangular distribution is √3. -LJFarrell
Note on G7: Columns F and G: If Column E is Type A, then the distribution of expected values is assumed to be "normal," also called Gaussian. If Column E is Type B, other distributions can be assumed. If a full range of possible values is estimated in column C , then a rectangular distribution of possible values is usually assumed. If so, then the standard uncertainty is equal to that full range divided by 2√3. If only one side of the range of possible values is shown in Column C, as in cells C9 and C10 here, then the divisor for a rectangular distribution is √3. -LJFarrell
Note on H7: Column H: The more data points you have, the more confident you can be that a statistical calculation of 95 % confidence intervals is accurate. The degrees of freedom may be used to estimate confidence intervals. For type B components calculated by non-statistical methods, people usually calculate a conservative worst case estimate of error and assume infinite d.o.f. -Chris M
Note on I7: Column I = Column C / Column G. This procedure converts the stated uncertainty values in Column C into standard uncertainties. Now we are in a position to combine apples with apples. The quadratic sum is the combined standard uncertainty, and it appears in cell I14 here. -LJFarrell
Note on J7: Column J: See the formulas in the cells in this column. The largest component contributes the most to the resulting uncertainty. -Chris M

## Shotgun
Measurement Uncertainty Estimation Form
Measurement: |  |  | Trigger Pull Force of Shotgun 1
Range of measurement values: |  |  | 3.77 lbsf - 5.92 lbsf
Procedure name and revision: |  |  | Example Uncertainty Budget
Estimation prepared by: |  |  | Task Group on Uncertainty of Measurement |  |  |  |  | Date Prepared: |  | 2019-07-04
Attention! See explanatory notes on row 7 and column A.
Line Item | Uncertainty Component | Value | Units | Type | Statistical Distribution | Divisor | Degrees Freedom (n-1) | Relative Standard Uncertainty (%) | Component Contribution %
1 | Measurement Process Reproducibility (from page 1) | 0.409 | lbsf | A | normal | 1 | 19 | 0.409 | 99.6292
2 | Calibration Certificate for Force Gauge - ± 1 % | 0.043 | lbsf | B | rectangular | 1.73 | ∞ | 0.024855 | 0.36795
3 | Thermal Dependence of Force Gauge. Assume ± 5°F uncertainty in temperature. | 0.0038 | lbsf | B | rectangular | 1.73 | ∞ | 0.0021965 | 0.0028735
4 | Other uncertainty sources, etc.
[8 empty] | Sum | 100
| Combined Standard Unc | uc | [5 empty] | 0.40976
| Expanded Unc | U (k=2) | [5 empty] | 0.81952
| Expanded Unc | U (k=3) | [5 empty] | 1.2293
| Reported Uncertainty: | 0.81952 |  | k=2
| Reported Uncertainty: | 1.2293 |  | k=3
NOTE: Regardless of the number of digits that are showing in a cell, Excel carries the maximum number of significant figures in the background and will use the entire number for further calculations
Note on A8: The first component of uncertainty is an estimate of the standard deviation of the measurements by the observers in the study. It is copied from cell T36 shown on sheet one. -Chris M
Note on A9: The force gauge should have an uncertainty obtainable from the vendor or from a calibration lab. For simplicity, we assume here a maximum uncertainty of ± 1 % with a rectangular distribution. This component of uncertainty may also be quoted by the vendor as a fixed amount (example - 0.1 lbsf). With a measured average value of 4.285 lbsf shown on page 1, the estimated maximum uncertainty is 0.0.043 lbsf. -Chris M
Note on A10: It seems that force measuring instruments should have some temperature dependence, particularly if they have any spring components. The last worksheet contains a rough estimate of possible error in a spring constant, yielding ± 0.088 % (outer limit of error) for a temperature uncertainty of ± 5 °F. With a measured value of 4.285 lbsf, the maximum uncertainty is ± 0.0038 lbsf. -Chris M
Note on A11: other sources, etc. -Chris M
Note on B7: Column B shows the different sources of uncertainty that are considered significant. -Chris M
Note on C7: Column C shows the estimated uncertainty values from each source. This is either derived from data showing variation of the source component or is derived from our understanding and modelling of the possible bias or variation. -LJFarrell
Note on C17: Limit to at most two significant figures unless there is a documented rationale for reporting additional significant figures. The laboratory procedure for estimating measurement uncertainty must include the process for rounding the expanded uncertainty. -LJFarrell
Note on D7: Column D: The source components in this template are all converted to lbsf. If one source uncertainty component has different units from the others, such as newtons or %, the values need to be converted into equivalent values in lbsf, and the converted values need to be recorded in column C for combination with the other uncertainty components. -LJFarrell
Note on E7: Column E: If the uncertainty value in column C is derived by statistical methods yielding, say, a standard deviation, then the uncertainty is Type A. If the value is derived by other means such as from manufacturer specifications or from a calculation of possible bias, then the uncertainty is Type B. -LJFarrell
Note on F7: Columns F and G: If Column E is Type A, then the distribution of expected values is assumed to be "normal," also called Gaussian. If Column E is Type B, other distributions can be assumed. If a full range of possible values is estimated in column C , then a rectangular distribution of possible values is usually assumed. If so, then the standard uncertainty is equal to that full range divided by 2√3. If only one side of the range of possible values is shown in Column C, as in cells C9 and C10 here, then the divisor for a rectangular distribution is √3. -LJFarrell
Note on G7: Columns F and G: If Column E is Type A, then the distribution of expected values is assumed to be "normal," also called Gaussian. If Column E is Type B, other distributions can be assumed. If a full range of possible values is estimated in column C , then a rectangular distribution of possible values is usually assumed. If so, then the standard uncertainty is equal to that full range divided by 2√3. If only one side of the range of possible values is shown in Column C, as in cells C9 and C10 here, then the divisor for a rectangular distribution is √3. -LJFarrell
Note on H7: Column H: The more data points you have, the more confident you can be that a statistical calculation of 95 % confidence intervals is accurate. The degrees of freedom may be used to estimate confidence intervals. For type B components calculated by non-statistical methods, people usually calculate a conservative worst case estimate of error and assume infinite d.o.f. -Chris M
Note on I7: Column I = Column C / Column G. This process converts the stated uncertainty values in Column C into standard uncertainties. Now we are in a position to combine apples with apples. The quadratic sum is the combined standard uncertainty, and it appears in cell I14 here. -LJFarrell
Note on J7: Column J: See the formulas in the cells in this column. The largest component contributes the most to the resulting uncertainty. -Chris M